## Slide 1: Alignment procedure for SPCM
Morgan Blevins
Aug 30, 2024

## Slide 2
Laser w/ 2-axis adjustable mount
Mirror #3
Beam splitter
Beam compensator
Mirror #2
Mirror #1
iris
iris
Chopper
Beam splitter
Beam compensator
iris
iris
ND filter
Mirror #2
Mirror #3
Laser w/ 2-axis adjustable mount
Chopper
Objective

## Slide 3
Adjust mirror #1 and #2 (or #1 and the laser holder) back and forth until the laser is aligned all along the cage section when the iris is moved back and forth
Mirror #2
Step 1
Note: I had to be pretty careful about how I locked down the mirror #2 to the vertical post. If its too high or low it bends the posts and make the system un-level.

## Slide 4
Adjust mirror #3 until the laser is centered in the iris
Step 2
Iris
Can also place a handing alignment marker here
Beam compensator is playing an important role in maintaining alignment to the objective.
Can also monitor the beam profile on the camera and see that is centered

## Slide 5: Beam compensation must be properly aligned to the optical path and the beamsplitter
The beamsplitter also has a marker showing the top
There is a marker showing the top of the compensator
The wedge angle has a certain orientation

## Slide 6: NOTE I’ve used a mounting base to tilt up the cage so it’s level.
I used the level app on my phone
In the future this should be done more exactly/precisely- should consult Thorlabs for best practice on this.
Mounting block